## Slide 1: Automated ICD Testing Tool
Marco Beccani, Kuk Jang, Houssam Abbas, Zhihao Jiang, and Rahul Mangharam
Department of Electrical and Systems Engineering
University of Pennsylvania
Jackson Liang, MD and Sanjay Dixit, MD
Hospital of the University of Pennsylvania

## Slide 2: ICD Teardown

## Slide 3
ICD Teardown
Memory
Telemetry
ASIC
DC-DC
converter
Beeper

## Slide 4
ICD Teardown
RF Antenna
Capacitor
Power Module
Notes: And these are the components

## Slide 5
High Voltage Output Switches
Pace Output Circuit (Pulse Generator)
DF +
(Anode)
DF -
(Cathode)
ring
tip
Filter and High Voltage Protection Circuit
(Analog and Digital)
Sense Amplifier
Pacing Circuitry
Adaptive Sensors
Custom ASIC
High Voltage Output Switches
Charging Circuitry
Low Power Microprocessor
RAM and ROM
ICD Functional Block diagram
Telemetry
Beeper
Heart Model

## Slide 6
ICD Hardware components
CPU/ASIC
Microprocessor
RF Link
System Management
BATTERY
5-8 y monitoring 0.63 Ah
5-8 y pacing 0.5 Ah
200 Shocks 0.37 Ah
TOTAL 1.5 Ah
CONNECTOR BLOCK
Suture Anchor
Antenna
Leads
CAPACITOR
C = 80-200 µF
V = 700 – 1k V
Energy 25-40 J
Notes: Talk about the components in details here is summarized
…
Now one of the main problem of ICDs is the fact that they deliver inappropriate shocks…. (click)

## Slide 7: Device Validation
Device in the loop testing
Open ICD Platform
Notes: In the next slides I will describe how our toolchain can be used to test a real ICD device as well our implementation of it, the Open ICD.

## Slide 8
Conformance Testing
Open ICD
Heart Model
Commercial ICD
Scenario 1
Scenario 2
Scenario 3
Device Programmer
Algorithms implemented: Sensing, Detection, Therapy
Device in the loop testing
Notes: with our hearth model we can generate scenarios to fed the ICD leads. This can be either a commercial device or our implementation of an ICD (OPEN ICD).
If the device delivers a therapy the toolchain can detect it. In this way we can compare the scenario fed to the commercial device and the open ICD and understand if a therapy decision is matching for both the devices.

## Slide 9
Heart Model
Commercial ICD
The scenario is fed to a commercial ICD
Atrial signal
Ventricular signal
Shock signal
Hardware in the loop testing: ICD
Notes: Now in the case of the commercial device, the three signals Atrial Ventricular and shock are fed to the device. The shock signal is also used as feedback for the toolchain.

## Slide 10
Hearth Model
Open ICD
The scenario is fed to our implementation of an ICD algorithm (Open ICD)
Data acquisition & Shock Generation
fsample = 1 kHz
DAQ – USB
Atrial signal
Ventricular signal
Shock signal
Hardware in the loop testing: Open ICD

## Slide 11: Toolchain

## Slide 12
16 BIT
DAC
1 KHz
Atrial
Ventricular
Shock
Amplifiers
Shock Feedback
ADC
Shock Protection Circuit
-7mV +7 mV
800 V – 30 J
A
V
S
Shock
Shock Feedback
Microprocessor
Analog Front End
System Block Diagram
SPI
USB

## Slide 13
Signal Pipeline: Normalized AAEL

## Slide 14
Signal Pipeline: DAC Output

## Slide 15: Attenuation Stage
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)

## Slide 16
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)
Attenuation Stage

## Slide 17: Analog Isolation Amplifier
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)

## Slide 18: Shifting and buffering
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)

## Slide 19
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)

## Slide 20
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)
Attenuation

## Slide 21
DAC
Attenuation Stage
(1/8)
Analog Isolation Amplifier ( X 8)
Signal Shifting ( ± 1 V ) and buffering
Attenuation Stage
( 2/3)
Attenuation

## Slide 22: Shock Signal detection
800 V, t = 2 ms

## Slide 23: Shock Protection and Attenuation
SHOCK PROTECTION Transient Voltage Suppressor (TVS) 8 V
Attenuation
Isolation
ADC

## Slide 24: PCB Layout

## Slide 25
MUX
DAC
Attenuation
Isolation
Shifting and buffering
Shock Protection
Attenuation

## Slide 26: Live Demo

## Slide 27: Scope Shock Protection Output
Shock Event

## Slide 28: Inappropriate Shock Scenario
Shock Event
Shock Event
Commercial Device
Commercial Device

## Slide 29: Inappropriate Shock Scenario: results
Shock Event
Open ICD Inhibit Output
Open ICD Shock Output

## Slide 30: Inappropriate Shock Scenario: results
The commercial device and our open ICD delivered a therapy for the inappropriate shock scenario

## Slide 31: Research Directions
Testing multiple devices at the same time
Knowledge of the discriminator not black box testing
Sensitivity testing
Patient specific device parameter setting